## Paper records
ID | Record type | What Phase 02 determined | Predicate rule the paper carries | Location stated | Individual named | Period derived | 211.180(d) reproduction | Medium | What would be needed
PP-01 | RT-09 Tablet hardness and friability test printout, in-process physical testing — held on SYS-062 Ridgemont and SYS-063 Cary | Not a Part 11 record. The printout route's three conditions, two stated and one derived, all hold, and they hold on facts. The ticket carries the individual measurements, the mean, the instrument identity, the date and time and the sample identity, which is the whole of what 21 CFR 211.188(b)(5) requires be recorded of this in-process control; the signed ticket is what the reviewer reads; and the testers retain no retrievable result and have no export capability — EV-223, AS-14. Phase 07 does not disturb this | 21 CFR 211.188(b)(5), read with 21 CFR 211.110(a) and (c). Tablet hardness and friability are not among the six in-process controls enumerated at 21 CFR 211.110(a)(1) to (6); they are in-process controls under the paragraph's non-exhaustive opener | No | No | No | No | Thermal | A stated retention location; a named individual accountable for the availability of the paper under 21 CFR 211.180(c); the period 21 CFR 211.180(a) already sets, read and recorded; and either a reproduction under 21 CFR 211.180(d) or a statement of the medium's archival life from a source Helix can name
PP-02 | RT-22 Analytical balance weighing printout — held on SYS-064 Ridgemont and SYS-065 Cary | Not a Part 11 record. The three conditions, two stated and one derived, all hold. The ticket carries the mass, the balance identity, the date and time and the analyst's identification, which is what 21 CFR 211.194(a)(3) requires; it is affixed to the worksheet and countersigned; and the balances hold no record after the ticket prints — EV-223, AS-14. Phase 07 does not disturb this | 21 CFR 211.194(a)(3) — a statement of the weight or measure of sample used for each test, where appropriate | No | No | No | No | Thermal | A stated retention location; a named individual accountable under 21 CFR 211.180(c); a derived period, and no predicate rule supplies one for this record type on its own — the ticket travels with the worksheet and nothing records what happens to it when the worksheet is filed; and either a reproduction under 21 CFR 211.180(d) or a statement of the medium's archival life
PP-03 | RT-23 Manual pH and conductivity measurement printout — held on SYS-066 Ridgemont and SYS-067 Cary | Not a Part 11 record. The three conditions, two stated and one derived, all hold. A single value with its temperature and calibration status; there is no underlying data set from which a different value could be rendered, so the record is static; the printout is attached to the worksheet and signed; and the bench meters have no storage and no export port — EV-223, AS-14. Phase 07 does not disturb this | 21 CFR 211.194(a)(4) — a complete record of all data secured in the course of each test, including all graphs, charts and spectra from laboratory instrumentation, properly identified to show the specific component, drug product container, closure, in-process material or drug product, and lot tested | No | No | No | No | Thermal | A stated retention location; a named individual accountable under 21 CFR 211.180(c); a derived period; and either a reproduction under 21 CFR 211.180(d) or a statement of the medium's archival life. PR-16 records the replacement risk: a bench meter that logs to a network share ends condition three the day it is installed
PP-04 | RT-29 Portable temperature calibrator certificate printout — held on SYS-068 Ridgemont | Not a Part 11 record. The three conditions, two stated and one derived, all hold. The certificate carries the as-found and as-left values, the reference standard, the tolerance, the date and the technician, which satisfies the written-record obligation at 21 CFR 211.68(a), read with the programme content required by 21 CFR 211.160(b)(4). Neither paragraph prescribes certificate content; condition one holds because the certificate is a fixed record of a single check with no underlying data set from which a different result could be rendered. It is filed in the calibration record, which is relied on, and the calibrators hold no retained certificate — EV-223, AS-14. Phase 07 does not disturb this | 21 CFR 211.68(a) — the written records of calibration checks and inspections — read with 21 CFR 211.160(b)(4), which sets the content of the calibration programme | No | No | No | No | Plain paper | A stated retention location; a named individual accountable under 21 CFR 211.180(c); a derived period; and a reproduction under 21 CFR 211.180(d) where the original is not retained. The medium is not thermal and the archival-life question is different in kind, not absent